## Slide 1: Sorter #1 HOOD ARIA 2V-4B-2R
Lasers
Detectors
RED
640
VIOLET
407
BLUE
488
BV421
PACBLUE
DAPI
BV510
AMCYAN
CFP
FITC
GFP
ALEXA488
PER-CY5.5
PE
DSRED
PETXRED
PI
PE-CY5
APC ALEXA647
APC-Y7
PE-CY7

## Slide 2: Sorter #2 Self Aria 2U-5V-2B-4Y-3R
Lasers
Detectors
AP-CY7
PE-CY7
BV750
RED
633
VIOLET 407
BLUE
488
Y/G
568
BUV396
DAPI
CFP
BV421
PACBLUE
HOESCHT
RED
BV510
AMCYAN
CFP
BV605
QDOT 605
BV711
FITC
GFP
ALEXA488
PER-CY5.5
A700
PE
DSRED
PETXRED
PI
MCHERRY
PE-CY5
APC ALEXA647
UV
355
CFP
449
BV660

## Slide 3: Sorter #3 CLS ARIA 2V-2B-4Y-2R
Lasers
Detectors
RED
640
VIOLET
407
BLUE
488
Y/G
568
BV421
PACBLUE
DAPI
BV510
AMCYAN
CFP
FITC
GFP
ALEXA488
PER-CY5.5
PE
DSRED
PETXRED
PI
MCHERRY
PE-CY5
APC ALEXA647
AP-CY7
PE-CY7

## Slide 4: Sorter #4 FUSION ARIA 3V-2B-4Y-2R
Lasers
Detectors
APC ALEXA647
RED
640
VIOLET
407
BLUE
488
Y/G
568
BV421
PACBLUE
DAPI
BV510
AMCYAN
CFP
FITC
GFP
ALEXA488
PER-CY5.5
PE
DSRED
PETXRED
PI
MCHERRY
PE-CY5
BV750
AP-CY7
PE-CY7

## Slide 5: Analyzer #1 LSR-ii 2U-6V-2B-4Y-3R
Lasers
Detectors
AP-CY7
PE-CY7
BV750
RED
640
VIOLET 407
BLUE
488
Y/G568
BUV396
DAPI
BV421
PACBLUE
BV510
AMCYAN
CFP
BV605
QDOT 605
BV711
FITC
GFP
ALEXA488
PER-CY5.5
A700
PE
DSRED
PETXRED
PI
MCHERRY
PE-CY5
UV 355
BV660
DAPI
HOESCHT
RED
APC ALEXA647

## Slide 6: Analyzer #2 FORTESSA 8th 6V-2B-4Y-3R
Lasers
Detectors
RED
640
VIOLET
407
BLUE
488
Y/G
568
BV421
PACBLUE
DAPI
BV510
AMCYAN
CFP
FITC
GFP
ALEXA488
PER-CY 5.5
PE
DSRED
PETXRED
PI
MCHERRY
PE-CY5
APC ALEXA647
BV750
AP-CY7
PE-CY7
BV605
QDOT 605
BV660
BV711
ALEXA700

## Slide 7: Analyzer #3 FORTESSA 6th 4U-5V-2B-4Y-3R
Lasers
Detectors
PE-CY7
AP-CY7
BV750
RED
640
VIOLET 407
BLUE
488
Y/G
568
BUV396
DAPI
BUV 661
BV421
PACBLUE
BUV 737
BUV496
BV510
AMCYAN
CFP
BV605
QDOT 605
BV711
FITC
GFP
ALEXA488
PER-CY5.5
A700
PE
DSRED
PETXRE
PI
MCHERRY
PE-CY5
APC
ALEXA647
UV 355

## Slide 8: Sorter #1 HOOD ARIA 2V-4B-2R
Technical Specifications
Quartz cuvette flow cell is gel-coupled by refraction index-matching optical gel to the florescent objective lens. Sort fractions collected jet-air
3 laser
11 parameters: 9 fluorescent and two scatter parameters + Time
Sheath pressure is adjustable from 20,35,70 PSI at normal operation
Maximum of 4 way sort fractions
Aerosol Management Option(AMO) with Bio-Bubble negative pressure
Bio Safety Level 2+ environment
262,144 channel(18bit) dynamic range on all parameters
Window xp-Pro OS, Diva software acquisition

## Slide 9: Sorter #2 Self Aria 2U-5V-2B-4Y-3R
Technical Specifications
Quartz cuvette flow cell is gel-coupled by refraction index-matching optical gel to the florescent objective lens. Sort fractions collected jet-air
5 laser-separate interrogation of sample for each laser
18 parameters: 16 fluorescent and two scatter parameters + Time
Sheath pressure is adjustable from 20,35,70 PSI at normal operation
Maximum of 4 way sort fractions
Sample collection cooling 4◦, 20 ◦, 37 ◦, and 42 ◦C(water re-circulator for refrigeration/heating)
Aerosol Management Option(AMO) with Bio-Bubble negative pressure
Bio Safety Level 2+ environment
262,144 channel(18bit) dynamic range on all parameters
Window xp-Pro OS, Diva software acquisition

## Slide 10: Sorter #3 CLS ARIA 2V-2B-4Y-2R
Technical Specifications
Quartz cuvette flow cell is gel-coupled by refraction index-matching optical gel to the florescent objective lens. Sort fractions collected jet-air
3 laser
11 parameters: 9 fluorescent and two scatter parameters + Time
Sheath pressure is adjustable from 20,35,70 PSI at normal operation
Maximum of 4 way sort fractions
Aerosol Management Option(AMO) with Bio-Bubble negative pressure
Bio Safety Level 2+ environment
262,144 channel(18bit) dynamic range on all parameters
Window xp-Pro OS, Diva software acquisition

## Slide 11: Sorter #4 FUSION ARIA 3V-2B-4Y-2R
Technical Specifications
Quartz cuvette flow cell is gel-coupled by refraction index-matching optical gel to the florescent objective lens. Sort fractions collected jet-air
4 laser
16 parameters: 14 fluorescent and two scatter parameters + Time
Sheath pressure is adjustable from 20,35,70 PSI at normal operation
Maximum of 4 way sort fractions
Aerosol Management Option(AMO) with Bio-Bubble negative pressure
Bio Safety Level 2+ environment
262,144 channel(18bit) dynamic range on all parameters
Window xp-Pro OS, Diva software acquisition

## Slide 12: Analyzer #1 LSR-ii 2U-6V-2B-4Y-3R
Technical Specifications
5 laser, separate interrogation of sample for each laser
17 color detection + 2 scatter parameters (from 488nm)
Hydrodynamically focused sample interrogated within a quartz cuvette flow cell
Sample aspiration rates at 12ul, 35ul, 60ul per minute
262,144 channel(18bit) dynamic range on all parameters
PC Window XP, Diva version 613

## Slide 13: Analyzer #2 FORTESSA 8th 6V-2B-4Y-3R
Technical Specifications
4 laser, separate interrogation of sample for each laser
11 color detection + 2 scatter parameters (from 488nm)
Hydrodynamically focused sample interrogated within a quartz cuvette flow cell
Sample aspiration rates at 12ul, 35ul, 60ul per minute
262,144 channel(18bit) dynamic range on all parameters
PC Window XP, Diva version 6.3

## Slide 14: Analyzer #3 FORTESSA 6th 4U-5V-2B-4Y-3R
Technical Specifications
5 laser, separate interrogation of sample for each laser
11 color detection + 2 scatter parameters (from 488nm)
Hydrodynamically focused sample interrogated within a quartz cuvette flow cell
Sample aspiration rates at 12ul, 35ul, 60ul per minute
262,144 channel(18bit) dynamic range on all parameters
PC Window XP, Diva version 613

## Slide 15
FLOW
CYTOMETRY
RESEARCH
LAB
F
L
O
W
C
Y
T
O
M
E
T
R
Y
R
E
S
E
A
R
C
H
L
A
B
O
R
A
T
O
R
Y

## Slide 16: Analyzer #3 FORTESSA 6th CONFIG
AP-CY7
PE-CY7
BV750
UV
RED
VIOLET
BLUE
Y/G
BUV396
BUV 661
BV 421
BUV 737
BUV496
BV510
BV605
BV711
FITC
PER-CY
5.5
A700
PE
PETXRED
PE-CY5
APC
Lasers
Detectors

## Slide 17: MACSVYB 2V-2B-4Y
Lasers
Detectors
VIOLET
407
BLUE
488
BV421
PACBLUE
DAPI
BV510
AMCYAN
CFP
FITC
GFP
ALEXA488
PETXRED
PI
APC
PE-CY 5
APC-CY7
PE-CY7
Y/G
568
PE
DSRED